Operation Manual

HMGTU-200kVA/500kV

AC Hipot Tester

WUHAN GOLDHOME HIPOT ELECTRICAL CO., LTD.

Tel：+86-18627130886                Fax：+86-27-84620387

Email: goldhome01@cablehipot.com        Web: www.hvcablehipot.com

Add: No. A25-1, Huading Mould Industrial Park, Hannan, Wuhan, Hubei, China

- Scope of application

The test set is suitable for factory testing of voltage levels below 500kV, cable insulation withstand voltage, and partial discharge measurement.

- Installation and usage conditions

- Altitude: No requirements for fully enclosed testing equipment

- Environmental temperature: -10℃~+40℃

- Relative humidity of air: No requirements for fully enclosed testing equipment

- Have a good grounding system.

- Installation and use location: indoor use

- The waveform of the power supply voltage is a sine wave, with waveform distortion ≤ 3%

- Independent power supply, three-phase five wire power supply, grounding, spatial partial discharge background<1PC

- There is a reliable grounding with a grounding resistance＜0.5Ω.

- Basic technical requirements

The complete testing set mainly consists of a control system, a high-voltage power supply system, a test sample interface system, a low-voltage power supply system. Conduct the test in an unshielded environment, with the harmonic component of the test power supply being less than 5%, and the local discharge of the entire system being less than 1pC at rated voltage. The annual leakage rate of each sealed chamber is less than 0.5%.

- Configuration of test set

- Operation desk(remote control unit)

Technical parameters

- Model: HMTC-200/500

- S/N: 2509052

- Input voltage: 220V

- Frequency: 50Hz

- Isolation/compensation/voltage regulation cabinet

Technical Parameters

- Model: HM200kVA

- S/N: 2509052

- Power Supply: Single phase, 380V/50Hz

- Input Capacity: 50kVA

- Input Current: 131.5A

- Output Capacity:200kVA

- Output Voltage: 1000V

- Output Current: 200A

- Power Insulation Transformer: 50kVA/400V/400V

- Electrical Voltage Regulator: 50kVA/400V

- Isolation Transformer: 200kVA/400V/1000V

- Compensation Reactor: 25kVA/400V*8

- Test Transformer

When it is necessary to open the flange, the air pressure in the flange area needs to be reduced to 0, and the air pressure in the transformer body does not need to be reduced.

The maximum pressure difference between the docking flange and the transformer body is 0.7MPa.

During normal testing, the tank pressure is maintained at 0.5±5%MPa.

Technical Parameters

- Model:HMGTU-200kVA/500kV

- S/N: 2509052

- Power Supply: AC 380V/50Hz, Single phase

- Rated Capacity: 200kVA

- Rated HV: 500kV

- Rated HV Current: 400mA

- Rated LV: 1000V

- Rated LV Current: 200A

- Rated Frequency: 50Hz

- Insulation Medium: SF6

- Inflation pressure: 0.5Mpa

- Waveform Distortion rate: ≤3%

- Impedance Voltage of test transformer:＜12%

- Capacitance of coupling capacitor: 300pF

- PD level: ≤1pC at rated

- Rated HV current:400mA

- Rated LV current:200A

- Metering voltage of test transformer: 200V

- Accuracy of voltage divider: 1.5%

- Wire Connection

- Operation desk

- The power switch and emergency button are on the left hand of the desk.

Computer

Power Strip

Aviation Connector

To the cabinet

Attention:

Please select the LAN port on the right, as the system presets parameters through this interface

- Isolation/compensation/voltage regulation cabinet

- Power switch on the right hand of the cabinet, it is the power switch for the entire testing system.

- Input A&B: Power supply of the entire system

- C&N: Connect with the AC mains

- Meter: connect with the meter of test transformer

- Divider signal: the HV output signal, connect with the voltage divider

- HV current: connect with the test transformer

- Aviation connector: connect with the operation desk

- Optical fiber: connect with the quick protection box

- Output 0~1000V: connect with the input of the test transformer

- PE: earthing

- Test Transformer

- Input: connect with ouput terminals of the cabinet

- Meter: connect with the meter of the cabinet

- Signal: this is the HV current sigbal, connect with the HV current of the cabinet

- PE: earth, at the bottom, picture as below

PE

- PE: at the base, bottom left corner, picture as below

PE

- HV voltage divider

- *Quick protection box: When the test sample breaks down, it provides quick protection for the test transformer.

PE

Divider Signal

Connect to Cabinet

Optical Fiber

Connect to Cabinet

- Attention: After the pulse calibration is complete, wrap the grounding wire around the pole as shown in the figure above.

- Coupling capacitor

Impedance Matching Unit (for PD signal acquisition)

PD Signal

Connect to

PD tester

PE

- Operation

- Icon of GTU Test and Control System, and main interface:

- “SIMATIC”

Note: The actual output voltage is based on the red rectangle “high voltage”

*The meter voltage is the output HV of the test transformer, and it connecting a capacitive voltage divider will have a series resonant rise effect, so the meter voltage will be lower than the actual HV voltage.

- Parameter setting of the system

- Test voltage: testing voltage you need

- Withstand Time: ≤10 minutes at rated current; Continuous operation for≤30 minutes at 2/3 rated current

- Boost speed: 100-3000V / S Maximum

- Step-down speed: 100-3000V/s Maximum

- Alarm voltage: when the testing voltage value is over the setted voltage, the main contactor will be turned off, the setted voltage setting value is 1.1 times of required testing voltage value

- Alarm current: when the low voltage current value is over the setted low voltage current value, the main contactor will be turn off, test stops.

- Flashover factor: Suggest 30% input, activate protection when voltage drops by more than 30%

- Auto Operation

- After completing the parameter settings, go to the fully automatic operation interface

- Start: Click Start, the device will automatically boost voltage, automatically start withstand voltage time, and the voltage will automatically drop to 0.

- Stop: Used when the experiment needs to be stopped

- Manual Operation

- After completing the parameter settings, go to the manual operation interface

- Start: start experiment

- Boost: voltage rises, single tap for fine-tuning, long press to maintain.

- Step-down: voltage falls, single tap for fine-tuning, long press to maintain.

- Timing: duration of the experiment

- Stop: end experiment

- VII-Safety Precautions

- 1. Set up the testing site according to relevant procedures, connect the equipment cables, and in areas where possible, have designated safety personnel on-site to provide guidance. Connect the grounding terminals of the operating table, transformer, and cabinet to the grounding grid, ensuring a reliable connection.

- 2. Connect the power cord and turn on the power switch. If you need to approach the test transformer, the test sample, or other connected equipment, please press the "Emergency Stop" button.

- 3. After power is connected, if the system is not in the zero position, it will automatically return to the zero position.

- 4. Do not place unrelated items on or around the equipment panel during the test.

- 5. Check the power supply voltage before powering on.

- 6. When replacing fuses and accessories, please use the same model as this instrument.

- 7. Protect this instrument from moisture and oil.

- 8. Before testing, ensure the tested equipment is powered off and disconnected from other live equipment.

- VIII-Transportation and Storage

- Transportation

- When transporting the equipment, it is recommended to use our company's instrument packaging wooden crates and shock-absorbing materials to avoid unnecessary damage and losses during transportation.

- When transporting equipment without wooden crates, stacking is not permitted.

- During transportation, the panel should face upwards.

- Storage

- The equipment should be placed in a dry, dust-free, well-ventilated room free of corrosive gases. Stacking is not permitted without wooden crate packaging.

- When storing the equipment, the panel should face upwards. Place moisture-proof materials under the equipment to prevent moisture absorption.

- IX-Unpacking and Inspection

- Unpacking Precautions

- Before unpacking, please ensure that the arrow mark on the outer packaging of the equipment is facing upwards. Do not strike the equipment forcefully when unpacking to avoid damaging it. Remove the equipment from the box, but retain the outer packaging and shock-absorbing materials. This facilitates future use during transportation and storage and also protects the environment.

- Inspection Items

- After unpacking, remove the equipment and check the equipment and accessories according to the packing list. If you find any shortages, please contact our company immediately, and we will provide you with services as soon as possible.